## Slide 1: Data acquisition and instrument communication
Innovative Teaching Approaches in development of Software Designed Instrumentation and its application in real-time systems
Faculty of Physics
Warsaw University of Technology
Zagreb University of
Applied Sciences
Faculty of Technical
Sciences
Ss. Cyril and Methodius
University
Faculty of Electrical Engineering
and Information Technologies
School of Electrical
Engineering
University of Belgrade

## Slide 2
2
Системи за аквизиција на податоци
Поглавја
Системи за аквизиција во LabVIEW
Аналогни влезови
Меморирање на податоци
Аналогни излези
Бројачи
Дигитални влезови/излези
Notes: This lesson introduces data acquisition.
This lesson covers:
A brief discussion of boards and DAQ concepts.
The organization of LabVIEW DAQ VIs.
How to acquire an analog reading.
How to output an analog signal.
How to use counters.
How to acquire data continuously.
The most important thing to remember about both Lesson 9 and 10 is that this course is not meant to be anything more than a BRIEF introduction to let them know this capability exists. If they are pressing for more information, point them to the other Customer Education courses.

## Slide 3
3
Преглед и конфигурација
Основна цел на системот за аквизиција е мерење на сигнали од реалното опгружување
Системите за аквизиција се состојат од:
Сетила или претворувачи
Обработка (кондиционирање)
PC-картичка за аквизиција
Драјвер
Софтвер
Notes: Before a computer-based measurement system can measure a physical signal, such as temperature, a sensor or transducer must convert the physical signal into an electrical one, such as voltage or current. You might consider the plug-in DAQ device to be the entire measurement system, but it is actually only one system component. You cannot always directly connect signals to a plug-in DAQ device. In these cases, you must use signal conditioning accessories to condition the signals before the plug-in DAQ device converts them to digital information. The software controls the DAQ system by acquiring the raw data, analyzing, and presenting the results.

## Slide 4
4
Хардверска конфигурација на системот
Measurement & Automation Explorer (MAX)
Notes: The Windows Configuration Manager keeps track of all the hardware installed in your computer, including National Instruments DAQ devices. If you have a Plug & Play (PnP) device, such as an E-Series MIO device, the Windows Configuration Manager automatically detects and configures the device. If you have a non-PnP device, or legacy device, you must configure the device manually using the Add New Hardware option in the Control Panel.
You can verify the Windows Configuration by accessing the Device Manager, available by selecting Start»Settings»Control Panel»System»Device Manager. You can see Data Acquisition Devices, which lists all DAQ devices installed in your computer. Double-click a DAQ device to display a dialog box with tabbed pages. The General tab displays overall information regarding the device. The Resources tab specifies the system resources to the device such as interrupt levels, DMA, and base address for software-configurable devices. The NI-DAQ Information tab specifies the bus type of the DAQ device. The Driver tab specifies the driver version and location for the DAQ device.
LabVIEW installs Measurement & Automation Explorer (MAX), which establishes all device and channel configuration parameters. After installing a DAQ device in your computer, you must run this configuration utility. MAX reads the information the Device Manager records in the Windows Registry and assigns a logical device number to each DAQ device. Use the device number to refer to the device in LabVIEW. Access MAX either by double-clicking its icon on the desktop or selecting Tools»Measurement & Automation Explorer in LabVIEW.

## Slide 5
5
Канали и задачи
Сигнали
ai 0
ai 1
ai 2
Sine Wave 1
Sine Wave 2
Sine Wave 3
Имиња на канали
Задачи
Timing and Triggering
Notes: In Traditional NI-DAQ you can configure a set of virtual channels, or a collection of property settings that include a physical channel, the type of measurement or generation specified in the channel name, and scaling information. In Traditional NI-DAQ and earlier versions, virtual channels are a simple method to remember which channels are used for different measurements. NI-DAQmx Channels are similar to the virtual channels of Traditional NI-DAQ.
NI-DAQmx also includes tasks which are integral to the API. A task is a collection of one or more channels and the timing, triggering, and other properties that apply to the task itself. A task represents a measurement or generation you want to perform.
Channels created only inside a task are local. Channels defined outside a task are global and can be used separately. Configuring virtual channels is optional in Traditional NI-DAQ and earlier versions, but is integral to every measurement you take in NI-DAQmx. In Traditional NI-DAQ, you configure virtual channels in MAX. In NI-DAQmx, you can configure virtual channels either in MAX or in a program, and you can configure channels as part of a task or separately.
In the example, analog input signals are routed into the DAQ system. The ai signals have been given channel names of Sine Wave 1, Sine Wave 2, and Sine Wave 3. The DAQ System is programmed to perform a task. A task is a collection of one or more channels, timing, and triggering. A task represents a measurement or generation you want to perform.

## Slide 6
Додатоци на системите за аквизиција
Notes: not in course manual:
Before the class starts Exercise, you should explain the following points about using the DAQ Signal Accessory (DSA):
The 5V line from the DAQ device supplies power to the DSA.
Analog input channels 1 and 2 are available at the quick connect terminal located in the lower central portion of the top panel. The IC temperature sensor is hard-wired to AI channel 0 and is located in the lower right hand corner of the top panel.
Analog output channels 0 and 1 are available at the quick connect terminal located at the lower central portion of the top panel.
The function generator is located in the lower central portion of the top panel. It produces a sine wave and a square wave. Use the Frequency Range selector switch to select a frequency range and then use the Frequency Adjust knob next to it to further adjust the frequency.
There is a row of four LEDs in the middle of the DAQ Signal Accessory that correspond to Lines 0 to 3 on Port 0 of the MIO board. The four LEDs use inverted logic.
Briefly describe the other parts of the DSA.
Instructor — make sure each student has at least one wire with stripped ends.

## Slide 7
7
Аквизиција во LabVIEW
NI-DAQ
Специфични VIs за следните задачи:
Аналогни влезови
Аналогни излези
Дигитални I/O
Операции за бројачи
NI-DAQmx
Драјвер за следна генерација:
VIs за специфични задачи
Една група VIs за секој тип мерења
Notes: The Data Acquisition palette in LabVIEW contains a palette for traditional NI-DAQ and one for NI-DAQmx.
Traditional VIs are divided by the type of measurement; DAQmx VIs are divided by the type of task.

## Slide 8
8
NI-DAQmx
Една група VIs за аналогни I/O, дигитални I/O, и операции за бројачи
DAQ Assistant Express VI
Брзо и едноставно програмирање на DAQ систем
Креирање логички задачи
Најголем број апликации можат да го употребуваат DAQ
Notes: The DAQmx-Data Acquisition palette contains all of the VIs necessary to perform analog I/O, digital I/O, and counter/timer operations. The VIs are organized so that the most common operations can be performed using the VIs. You can configure a task to perform a very specific function by using the Property Nodes in the palette. Many applications that do not require advanced timing and synchronization can be performed by using the DAQ Assistant Express VI. This course describes the use of the DAQ Assistant Express VI to perform data acquisition.

## Slide 9
9
NI-DAQmx Data Acquisition типови задачи
Мерењето се однесува на:
Аналоген влез
Аналоген излез
Влез за бројач
Излез за бројач
Дигитален I/O
Notes: The DAQ Assistant Express VI allows you to easily configure the data acquisition device. When you place the DAQ Assistant Express VI on the block diagram, a dialog box appears where you configure a local task to perform a specific measurement function. Creating a local task allows you to specify the exact type of measurement to take.
Once you create a task, the information for the local task is stored in the DAQ Assistant Express VI.
You can reconfigure the DAQ Assistant Express VI by double-clicking the VI and creating a new task.

## Slide 10
10
Аналоген влез
Notes: An analog input measurement type for a task can be voltage, temperature, strain, current, resistance, or frequency.
Each measurement type has its own characteristics, such as resistor values for current measurements, or strain gauge parameters for strain measurements.

## Slide 11
11
Аналоген влез, време и тригерирање
Дефинира број на примероци и фреквенција на земање одбироци
Дефинира старт и тип на тригер за одредена задача
Notes: Task Timing
When performing analog input, the task can be timed to Acquire 1 Sample, Acquire n Samples, or Acquire Continuously.
Acquire 1 Sample: acquires one value from an input channel and immediately returns the value. This operation does not require any buffering or hardware timing.
Acquire n Samples: One way to acquire multiple samples for one or more channels is to acquire single samples in a repetitive manner. However, acquiring a single data sample on one or more channels over and over is inefficient and time consuming. Moreover, you do not have accurate control over the time between each sample or channel. Instead you can use hardware timing, which uses a buffer in computer memory, to acquire data more efficiently.
Acquire Continuously: view, process, or log a subset of the samples as they are acquired.
Task Triggering
When a device controlled by NI-DAQmx does something, it performs an action. Two very common actions are producing a sample and starting a waveform acquisition. Every NI-DAQmx action needs a stimulus or cause. When the stimulus occurs, the action is performed. Causes for actions are called triggers.
Start Trigger: This trigger starts the acquisition.
Reference Trigger: This trigger establishes the reference point in a set of input samples. Data acquired up to the reference point is pretrigger data. Data acquired after the reference point is posttrigger data.

## Slide 12
12
Аналоген излез
Notes: Use analog output to perform digital-to-analog (D/A) conversions. An analog output task can be voltage or current.
To perform a voltage or current task, a compatible device must be installed that can generate that form of signal.

## Slide 13
13
Аналоген излез, време и тригерирање
Дефинира број на примероци и фреквенција на генерирање на примероците
Дефинира старт и тип на тригер за задачата
Notes: Task Timing
When performing analog output, the task can be timed to Generate 1 Sample, Generate n Samples, or Generate Continuously.
Generate 1 Sample: Use single updates if the signal level is more important than the generation rate. You can use software timing to control when your device generates a signal. This operation does not require any buffering or hardware timing.
Generate n Samples: You can use software timing or hardware timing to control when a signal is generated. With software timing, the rate at which the samples are generated is determined by the software and operating system instead of by the measurement device. With hardware timing, a TTL signal, such as a clock on your device, controls the rate of generation. A hardware clock can run much faster than a software loop. A hardware clock is also more accurate than a software loop. Use Generate n Samples if you want to generate a finite time-varying signal, such as an AC sine wave.
Generate Continuously: Continuous generation is similar to generate n samples, except that an event must occur to stop the generation.
Task Triggering
Start Trigger: This trigger starts the generation.
Reference Trigger: The reference trigger is not supported for analog output tasks.

## Slide 14
14
Бројачи
Бројачот е дигитален модул
Примена на бројачите:
Регистрирање настани
Мерење фреквенција
Мерење периода
Мерење позиција
Генерирање импулси
Gate
Source
Output
Count Register
Count register – Го меморира бројот на избројани импулси од страна на бројачот
Source – Влезот ја зголемува состојбата на бројачот
Gate – Влез за овозможи/оневозможи функција
Output – Геерирање на еден или серија импулси
Notes: A counter is a digital timing device. You typically use counters for event counting, frequency measurement, period measurement, position measurement, and pulse generation. A counter contains four main components: a count register, a source, a gate, and an output.
Count Register—Stores the current count of the counter. You can query the count register with software.
Source—An input signal that can change the current count stored in the count register. The counter looks for rising or falling edges on the source signal. Whether a rising or falling edge changes the count is software selectable. The type of edge selected is referred to as the active edge of the signal. When an active edge is received on the source signal the count changes. Whether an active edge increments or decrements the current count is also software selectable.
Gate—An input signal that determines if an active edge on the source will change the count. Counting can occur when the gate is high, low, or between various combinations of rising and falling edges. Gate settings are made in software.
Output—An output signal that generates pulses or a series of pulses, otherwise known as a pulse train.
When you configure a counter for simple event counting, the counter increments when an active edge is received on the source.

## Slide 15
15
Дигитални влезови/излези
Digital I/O служи за исчитување или запишување на една дигитална линија или цела порта
Дигитална порта е збир на дигитални линии
Notes: Measuring and generating digital values is used in a variety of applications including controlling relays and monitoring alarm states. Generally, measuring and generating digital values is used in laboratory testing, production testing, and industrial process monitoring and control.
Digital I/O can read from or write to a line or an entire digital port, which is a collection of lines.
You can use the digital lines in your DAQ device to acquire a digital value. This acquisition is based on software timing. On some devices, you can configure the lines individually to either measure or generate digital samples. Each line corresponds to a channel in your task.
You can use the digital port(s) in your DAQ device to acquire a digital value from a collection of digital lines. This acquisition is based on software timing. You can configure the ports individually to either measure or generate digital samples. Each port corresponds to a channel in your task.

## Slide 16
16
Заклучок
MAX е основна алатка затестирање и нагодување на системите за аквизиција во LabVIEW.
DAQ Assistant се користи за нагодување на системот за аквизиција и извршување на аквизициските задачи.
Најголем број апликации можат да го употребатDAQ Assistant-от. Во случаи каде потребно е напредно синхронизирање се користи NI-DAQmx.
DAQ Assistant се користи за читање/пишување од следните порти: аналоген влез/излез, дигитален влез/излез и операции со бројачи.
Notes: Do not immediately display this slide.
Suggested questions for class participation:
How are DAQ Boards configured?
How can you test the configuration of your board?
Point out that more in-depth DAQ information is available in the LabVIEW DAQ course.

## Slide 17
17
Контрола на инструменти
ПОГЛАВЈА
Преглед на контрола на инструментите
GPIB комуникација и нагодување
Instrument I/O Assistant
Virtual Instrument Software Architecture (VISA)
Инструментациски драјвери
Сериска комуникација
Notes: This chapter introduces Serial and GPIB instrument control using LabVIEW.
This lesson covers:
GPIB (IEEE 488) background.
Instrument drivers in LabVIEW.
GPIB control using VISA
Serial communication background.
Serial control using VISA
ASCII and binary waveform transfer.

## Slide 18
18
Преглед на контрола на инструментите
За да се контролираат инструментите треба да се познава следното:
Тип на конектор на инструментот − Тип на потребни кабли
Електрични карактеристики − Комуникациски протоколи
Драјвери на располагање
Инструменти
Компјутер
Notes: You are not limited to the type of instrument that you control if you choose industry-standard control technologies. You can use instruments from many different categories, including serial, GPIB, VXI, PXI, computer-based instruments, Ethernet, SCSI, CAMAC, and parallel port devices. This lesson describes the two most common instrument communication methods, GPIB and serial port communication.
Take note of the following with PC control of instrumentation:
Type of connector (pinouts) on the instrument
Cables needed—null-modem, number of pins, male/female
Electrical properties involved—signal levels, grounding, cable length restrictions
Communication protocols used—ASCII commands, binary commands, data format
Software drivers available

## Slide 19
GPIB комуникација
GPIB инструменти
GPIB
кабел
GPIB интерфејс
Notes: GPIB Communications
Briefly describe a typical system containing several GPIB devices connected to a computer. Describe that one device is the System Controller and the others are Talkers/Listeners. Each device on the GPIB has a unique address.
GPIB instruments offer test and manufacturing engineers the widest selection of vendors and instruments for general-purpose to specialized vertical market test applications. GPIB instruments have traditionally been used as standalone benchtop instruments where measurements are taken by hand. Many GPIB instruments used today are still used in this fashion.

## Slide 20
20
Запознавање со стандарди
1965
1987
1990
1992
1993
1975
HP designs HP-IB (Hewlett Packard Interface Bus)
HP-IB becomes IEEE 488
IEEE 488.2 adopted
IEEE 488 becomes IEEE 488.1
SCPI (Standard Commands for Programmable
Instruments) added to IEEE 488.2
IEEE 488.2 revised
HS488 proposed
1999
HS488 approved
Notes: Introduction of Standards to Instrument Communication
In 1965, Hewlett-Packard designed the Hewlett-Packard Interface Bus (HP-IB) to connect their line of programmable instruments to their computers. Because of its high transfer rates (nominally 1 MB/s), this interface bus quickly gained popularity. It was later accepted as IEEE Standard 488-1975. Today, the name General Purpose Interface Bus (GPIB) is more widely used than HP-IB.
Because the original IEEE 488 document contained no guidelines for a preferred syntax and format conventions, work continued on the specification to enhance system compatibility and configurability among test systems. This work resulted in a supplement standard IEEE 488.2, Codes, Formats, Protocols, and Common Commands, for use with IEEE 488 (renamed to ANSI/IEEE Standard 488.1-1987). IEEE 488.2 does not replace IEEE 488.1. Many devices still conform only to IEEE 488.1. IEEE 488.2 builds on IEEE 488.1 by defining a minimum set of device interface capabilities, a common set of data codes and formats, a device message protocol, a generic set of common device commands, and a new status reporting model.
In 1990, the IEEE 488.2 specification included the Standard Commands for Programmable Instrumentation (SCPI) document. SCPI defines specific commands that each instrument class, which usually includes instruments from various vendors, must obey. Thus, SCPI guarantees complete system compatibility and configurability among these instruments. It is no longer necessary to learn a different command set for each instrument in an SCPI-compliant system, and it is easy to replace an instrument from one vendor with an instrument from another.

## Slide 21
21
GPIB спецификации на хардверот
Максимална должина на каблите = 20 m
Макс број на инструменти = 15 (2/3 вклучени)
1
12
13
24
DIO5
DIO6
DIO7
DIO8
REN
GND (TW PAIR W/DAV)
GND (TW PAIR W/NRFD)
GND (TW PAIR W/NDAC)
GND (TW PAIR W/IFC)
GND (TW PAIR W/SRQ)
GND (TW PAIR W/ATN)
SIGNAL GROUND
DIO1
DIO2
DIO3
DIO4
EOI
DAV
NRFD
NDAC
IFC
SRQ
ATN
SHIELD
Notes: Hardware Specifications
The GPIB is a digital, 24-conductor parallel bus. It consists of eight data lines (DIO 1-8), five bus management lines (EOI, IFC, SRQ, ATN, REN), three handshake lines (DAV, NRFD, NDAC), and eight ground lines. The GPIB uses an eight-bit parallel, byte-serial, asynchronous data transfer scheme. This means that whole bytes are sequentially handshaked across the bus at a speed that the slowest participant in the transfer determines. Because the unit of data on the GPIB is a byte (eight bits), the messages transferred are frequently encoded as ASCII character strings.
Additional electrical specifications allow data to be transferred across the GPIB at the maximum rate of 1 MB/sec because the GPIB is a transmission line system. These specifications are:
A maximum separation of 4 m between any two devices and an average separation of 2 m over the entire bus.
A maximum cable length of 20 m.
A maximum of 15 devices connected to each bus with at least two-thirds of the devices powered on.
If you exceed any of these limits, you can use additional hardware to extend the bus cable lengths or expand the number of devices allowed.
Note For more information about GPIB, visit the National Instruments GPIB support website at http://www.ni.com/support/gpibsupp.htm.

## Slide 22
GPIB архитектура на софтверот— Windows
Driver
Software
(*.DLL)
Instrument
Driver VIs
NI-VISA
LabVIEW
Диагност. алатки:
GPIB Diagnostic Tool
NI-Spy
Алатки за конф.:
Measurement &
Automation Explorer
Interface Board
(GPIB, PXI,
VXI, computer based, etc.)
Notes: Underneath and between the different hardware components are the software components that are the heart of your instrumentation system. The figure above shows the software architecture for Windows 2000/XP from the perspective of the computer.
The instrument is connected to the computer through a built-in connector such as RS-232 or an installed interface board such as GPIB or VXI/MXI. Driver-level software for instrument is in the form of a Dynamic Link Library (DLL). In the example shown in the figure above, LabVIEW is the high-level application software layer. LabVIEW includes built-in functions for GPIB, serial, VXI, and computer-based instruments that load and use the installed driver software.
As with all Windows device drivers, any DLLs you install for interface boards will also interact with the Windows Registry so that resources such as base addresses, interrupt levels, and DMA channels can be assigned. Configuration and diagnostic software tools are also available with National Instruments hardware.
Refer in-depth questions to technical support or suggest that the students purchase the GPIB course.
Instructor - Show how the Measurement & Automation Explorer allows you to obtain information about your GPIB interfaces as well as the DAQ board.

## Slide 23
23
Нагодување на GPIB инструментите
Measurement & Automation Explorer (MAX)
Notes: Explain that the Measurement and Automation Explorer (MAX) allows you to:
Configure the GPIB card and search for instruments
Set the VISA alias name
Run NI-Spy
Configure IVI instrument drivers (be careful you may confuse some people if you go into a long dialogue about IVI).
If a GPIB+ card is installed you can run the GPIB analyzer software from MAX

## Slide 24
24
Што е Instrument I/O Assistant?
Пристап низ LabVIEW Express VI
Се користи а нагодување на комуникацијата и податоците низ чекор по чекор постапка
Notes: The Instrument I/O Assistant located on the Functions»Input and Functions»All Functions»Instrument I/O palettes is a LabVIEW Express VI that allows you to easily test communication with your instrument and develop a sequence of query, parse, and write steps. These steps can be saved as an Express VI for instant use or can be converted to a LabVIEW subVI. Use the Instrument I/O Assistant when an instrument driver is not available.

## Slide 25
25
Комуникација со инструмент
Notes: Communicating with an Instrument
Once you have placed the I/O Assistant on the block diagram, the wizard opens. The wizard starts in the Select Instrument step, where you can choose a GPIB or serial instrument. You also can check the interface properties from this window.
After selecting the instrument, you can add sequences to Query and Parse, Write, or Read and Parse. In addition, after you have set up a communication sequence with one instrument, you can set up additional instruments in the same Express VI.
In the figure above, a GPIB instrument was set up, then a query (*IDN?) was sent to the instrument. The response was automatically parsed, resulting in a string output.

## Slide 26
26
Инструментациски драјвери
Повеќе од 1200 LabVIEW инструментациски драјв.
Notes: LabVIEW provides more than 1200 LabVIEW instrument drivers from more than 50 vendors. You can use these instrument drivers to build complete systems quickly. Instrument drivers drastically reduce software development costs because developers do not need to spend time programming their instruments. You can reuse the drivers in a variety of systems and configurations.
LabVIEW instrument drivers simplify instrument programming to high-level commands, so you do not need to learn the low-level instrument-specific syntax needed to control your instruments. The programming is broken down into general functions for the DMM
These instrument drivers are called LabVIEW instrument drivers because the source code is graphical programming made from standard LabVIEW functions and VIs. LabVIEW instrument drivers are in the form of VI libraries and are organized into categories of instrument functions. You will learn about each category.

## Slide 27
27
Инсталирање и наоѓање на драјвери
Достапни на ni.com/idnet
Инсталирање во библиотека: LabVIEW 7.0\instr.lib directory
Пристап од палетата Functions»Input»Instrument Drivers
Notes: Installing and Accessing Instrument Drivers
You can find the LabVIEW instrument drivers on your LabVIEW installation CD or download them from the National Instruments web page at www.ni.com/idnet. The LabVIEW Instrument Wizard can automatically install instrument drivers.
If you install the LabVIEW instrument drivers from the CD yourself or download them from the web page, you first decompress the instrument driver files to get a directory of instrument driver files. Place this directory into the LabVIEW\instr.lib directory on your computer. The next time you launch LabVIEW, you can access the instrument driver VIs from the Input »Instrument Drivers subpalette of the Functions palette as shown in the slide above.

## Slide 28
28
Модел на драјверите
Notes: All instrument drivers in the library have the same VI Tree structure. Therefore, once you learn to use one instrument driver, all others have the same basic hierarchy. In fact, this hierarchy, sequence of VIs, and error checking are used in many other areas of I/O in LabVIEW such as file I/O, data acquisition (DAQ), and TCP/IP communications.
The structure of an instrument driver is shown in the figure above. The high-level functions are built from the lower-level functions. For the most control over the instrument, you would use the lower-level functions. However, the high-level functions like the Getting Started VI you used in the previous lesson are easy to use and have soft front panels that resemble the instrument. The component functions of instrument drivers are organized into the following categories:
Initialize
Configure
Action/Status
Data
Utility
Close

## Slide 29
29
Initialize • • •
Configure • • • • • • • •
Action/Status • • • • • • • • • • • •
Data • • • • • • • • • • • • • • • • • • • • • • • • • •
Utility • • • • • • • • • • • • • • • • • • • • • • • • • • • • • • •
Close • • • • • • • • • • • • • • • • • • • • • • • • • • • • • • • • • • • • • •
Vis за инструментациски драјвери
Notes: Instrument Driver VIs
Initialize—Initializes the communication channel to the instrument. The initialize VI can optionally perform an identification query and reset operation. In addition, it can perform any necessary actions to place the instrument in its default power-on state or other specified state.
Configure—A collection of VIs that configure the instrument for the operations you want to perform. An example is a function to set up the trigger rate.
Action/Status—Contains both action and status VIs. Action VIs cause the instrument to initiate or terminate test and measurement operations. Status VIs obtain the current status of the instrument or the status of pending operations. An example of an action function is Acquire Single Shot. An example of a status function is Query Transfer Pending.
Data—VIs that transfer data to or from the instrument such as reading a measured waveform from the instrument or downloading a waveform to the instrument.
Utility—VIs that perform a wide variety of useful functions such as reset, self-test, error query, and revision query.
Close—Terminates the communication channel to the instrument and de-allocates the resources set aside for that instrument.

## Slide 30
30
Сериска комуникација
Комуникација помеѓу PC и периферен уред
Податоците се испраќаат бит-по-бит
Се применува за мали податочни брзини и долги растојанија
RS-232 Instrument
RS-232 Cable
PC Serial
Port
Notes: Serial Port Communication
Serial communication is a popular means of transmitting data between a computer and a peripheral device such as a programmable instrument or even another computer. Serial communication uses a transmitter to send data, one bit at a time, over a single communication line to a receiver. You can use this method when data transfer rates are low or you must transfer data over long distances. Serial communication is popular because most computers have one or more serial ports, so no extra hardware is needed other than a cable to connect your instrument to the computer (or two computers together).
Serial communication requires that you specify four parameters: the baud rate of the transmission, the number of data bits encoding a character, the sense of the optional parity bit, and the number of stop bits. Each transmitted character is packaged in a character frame that consists of a single start bit followed by the data bits. A typical character frame encoding the letter “m” is shown in the slide above. Baud rate is a measure of how fast data is moving between instruments that use serial communication. RS-232 uses only two voltage states, called MARK and SPACE. In such a two-state coding scheme, the baud rate is identical to the maximum number of bits of information, including “control” bits, that are transmitted per second. MARK is a negative voltage and SPACE is positive; the slide above shows how the idealized signal looks on an oscilloscope. The truth table for RS-232 is:
Signal > +3 V = 0 Signal < -3 V = 1
The output signal level usually swings between +12 V and -12 V. The “dead area” between +3 V and -3 V is designed to absorb line noise.

## Slide 31
31
Сериско хардверско поврзување
RS-232
  DCE or DTE configurations
  9-pin or 25-pin
RS-422
  DCE or DTE
  8-pin
RS-485
  Multidrop
Pin DTE DCE
1 DCD Input Output
2 RxD I O
3 TxD O I
4 DTR O I
5 Com - -
6 DSR I O
7 RTS O I
8 CTS I O
9 RI I O
Notes: There are several recommended standards (RS) for serial communication. Each varies in hardware and software specifications. The students must be familiar with their instrument and what connector is used before you can begin controlling that device with their computer.
There are three main Serial I/O types:
RS-232—Connector found on most PCs. This is a single-ended communication method where only one device can be connected per port. Two connector types, 8- or 25-pin. Two configurations, DCE or DTE.
RS-422—Connector found on most Macs. This is a differential communication method. Connector has 8 pins.

## Slide 32
32
Сериска комуникација
Терминологија
Baud rate – бити во секунда
Data bits – инвертирана логика, LSB е прв
Parity – дополнителен бит за проверка на грешка
Stop bits – 1, 1.5, или 2 инвертирани бити закрај
Flow control – опции за хардверско и софтверско ракување
Notes: Instructor: See how many students are going to use or are interested in Serial I/O. Then cover the material in this and the next four slides accordingly.
Go over the following Serial I/O terminology:
Baud rate—How many bits per second are transferred on the serial cable.
Data bits—How many bits represent a data value.
Parity—Optional error checking bit that is added to the data.
Stop bits—Certain number of bits added to the end of each data transfer.
Flow control—Optional hardware or software handshaking parameters for communicating with a device.

## Slide 33
33
Примена на Instrument I/O Assistant за сериска комуникација
Се избираCOMX за адреса на инструментот
Се употребува I/O Assistant исто како кај GPIB
Notes: Use the I/O Assistant just like with GPIB communication.

## Slide 34
Innovative Teaching Approaches in development of Software Designed Instrumentation and its application in real-time systems
ITASDI
Co-funded by the
Erasmus+ Programme
of the European Union